## Comments
sl no | updated date | review status
1 | 2022-04-04 | ok | The BG95 MAIN UART CTS and RTS connection seems incorrect, please re-check and correct it.
2 | 2022-04-04 | ok | STM32L4 Debug UART TX and RX pins to the TC2030-NL is not crossed over. Match the pinout to --> https://www.tag-connect.com/product/tc2030-nl-ftdi-ttl-232rg-vip-usb-to-tc2030-no-legs-serial-cable
3 | 2022-04-04 | ok | RV3028 is not required, RTC_INT signal to to deleted, STM32L4 RTC is to be used
|  | ok | STM32L4 RTC to be supported in VBAT mode
|  | ok | Low Speed Clock external (LSE) needs to be connected to STM32L4 chip for using RTC in VBAT mode
As per datasheet section 3.33
"The RTC is functional in VBAT mode and in all low-power modes when it is clocked by the LSE. When clocked by the LSI, the RTC is not functional in VBAT mode, but is functional in all low-power modes except Shutdown mode."
4 | 2022-04-04 | ok | "BLDC_POWER_CTRL" to be connected to a DAC Channel
Since there is a shortage of pins, you can remove the CTS/RTS from the STSPIN32F0A chip
5 | 2022-04-04 | ok | Please check all your UART Tx, Rx, CTS and RTS connections. They do not seem to be crossed over correctly
I found issues with BG95, STSPIN220F0A but please re-check all of them
The connection should be like below
6 | 2022-04-04 | incorrect TXD & RXD are reversed | ESP32 Debug/PROG connection pinout is incorrect. It should match the pinout of the debug tool ESP-PROG. Check the link below
|  | see the ESP-PROG schematics | https://docs.espressif.com/projects/espressif-esp-iot-solution/en/latest/hw-reference/ESP-Prog_guide.html#communication-interface
7 | 2022-04-04 | ok | Do not connect KEY_IO_x signals to IO_Expander as it will be difficult to support the rotary encoder functionality on software
Consider moving all "CONN_x" signals to the IO Expander
There are many ENs, RSTs and GPIOs that can be moved to the expander chip
8 | 2022-04-04 | ok | Add a connection from the IO_Expander RESET signals to an MCU pin as done on C201 design
9 | 2022-04-04 | not answered | What will be the overall load on the 3.3V power supply? |  |  |  | total current 1.5A
|  |  | Can the LM27342SDX/NOPB handle the load? Note, apart from the on board chips it has to also power several (10 to 12) external sensors |  |  |  | LM27342SDX/NOPB max current 2A
Would it be better to have 2 separate 3.3V power supplies?
10 | 2022-04-04 |  | Bring the S201, S202 and Lidar schematics in to the sensors page
11 | 2022-04-04 | ok | The I2C_SYS bus has 11 slave devices now. To balance the load on the bus I think we should add another I2C bus (I2C_4)
If you free up PB10 and PB11, you can easily add the I2C_4 bus
I suggest moving the below devices to the new I2C bus
|  |  |  | CTP | 7BIT_I2C_ADDR: 0x38
|  |  |  | Lidar | 7BIT_I2C_ADDR: 0x52
|  |  |  | TS3USBCA410IRSVT | 7BIT_I2C_ADDR: 0x5D
12 | 2022-04-04 | could be made better | The schematic has become very complex. Can you make the MCU schematic doc page more readable by grouping signals as done on C201 and C204?
13 | 2022-04-06 | not done | The S201 & S202 connectors and accompanying circuitry are not needed. Please remove (or mask) them
14 | 2022-04-06 | not done | The rotary encoder connector is not required. Please remove it
15 | 2022-04-06 | not done | add another pin on the membrane keypad connector for the on/off button
16 | 2022-04-06 | not done | add 1 UVC LED on the main board. The LED connection to be made parallel to any one of the UVC FFC connector
also I think it is better to increase the value and wattage of the series resistors to the UVC LEDs
17 | 2022-04-07 | not done | Change stepper driver ribbon cable connector to Molex 5023821070 |  |  | replaced by 5023821270
18 | 2022-04-07 | not done | Change heating element connectors to
25W water heater connected by Molex 2121340002
5W heating coil connected by JST B2B-ZR-SM4-TF
19 | 2022-04-08 |  | TPS55340 will not work in the given input voltage range. Please consider using a different chip. MAX25240AFFF/VY+ can also be considered

## I2C_Buses
I2C_TS |  | I2C_HS |  | I2C_SYS |  | I2C_SYS_2
Temperature Sensor | 7BIT_I2C_ADDR: 0x76 | Humidity sensor | 7BIT_I2C_ADDR: 0x76 | BQ25792RQMR | 7BIT_I2C_ADDR: 0x6B | CTP | 7BIT_I2C_ADDR: 0x38
| 7BIT_I2C_ADDR: 0x77 |  | 7BIT_I2C_ADDR: 0x77 | LTC2944IDD#PBF | 7BIT_I2C_ADDR: 0x64 | Lidar | 7BIT_I2C_ADDR: 0x52
|  |  |  | PN5120A0HN1/C2,151 | 7BIT_I2C_ADDR: 0x2F | TS3USBCA410IRSVT | 7BIT_I2C_ADDR: 0x5D
|  | NDIR (MCU) | 7BIT_I2C_ADDR: ??? | PCAL9539AHF,128 main board | 7BIT_I2C_ADDR: 0x76
|  |  |  | BMA253 | 7BIT_I2C_ADDR: 0x19
|  |  |  | PCAL9539AHF,128 LTE board | 7BIT_I2C_ADDR: 0x74
|  |  |  | CMX655D | 7BIT_I2C_ADDR: 0x54
|  |  |  | STSPIN32F0A | 7BIT_I2C_ADDR: configurable

## UARTs
|  | TXD | RXD | CTS | RTS
USART1 | LTE | PG9 | PG10 | PG11 | PG12
USART2 | BLDC | PA2 | PA3 | PD3 | PD4
USART3 | ESP32 | PC4 | PC5 | PA6 | PB1
UART4 | PD Controller | PA0 | PA1
UART5 | C204 | PD2 | PC12

## Power
3.3V |  | max curr (mA) @ 3.3V | 6V | max curr (mA) @ 6V | 3.3V_BAT | 4V
STM32G0 |  | 6.7 | 3 x UVC | 90 | STM32L4S7 RTC | LTE BG95
TCPP01-M12 |  | 2 | STSPIN220 | 350 | MIC1555
TCPP02-M18 |  | 2.7 | Stepper driver | 350 | SIT1533
STM32L4S7 |  | 23
PN512 |  | 100
TFT |  | 70 |  | 790
TFT CTP |  | 10
MT29 |  | 35 | LM27342
SD Card |  | 200 | eff @ 60 deg | 87%
TS3USBCA |  | 0.04 | Pout | 6
PCAL9539A |  | 0.125 | Pin | 6.8966
4 x sensor connectors
| 4 x BMA280 | 0.52
| 2 x T+H | 0.26
| 2 x NDIR | 150
BMA253 |  | 0.13
Lidar |  | 24
LTE TS3USB3031 |  | 0.04
LTE PCAL9539A |  | 0.125
LTE TXB0108
ESP32 C3 Mini |  | 300
TPS61169
LDLN025
924.64
LM27342
| eff @ 60 deg | 72%
| Pout | 3.96
| Pin | 5.5